| There must be a calibration programme, described in a sufficiently detailed written procedure(s). The programme must include the following: - Mechanism, such as an Equipment Master List, for differentiating product critical from process /system, safety/environmental non-critical equipment and instrumentation. - Requirement that equipment/instrumentation be calibrated prior to initial use. - System for identifying equipment (e.g., ID number) and location and function label (TAG number) - Mechanism for establishing calibration tolerances taking into consideration the allowable or existing process tolerance and the capability of the equipment being calibrated. - System for assigning calibration frequencies. This should include the time window within which the calibration must be performed (e.g., every 6 months  2 weeks). Frequencies may be documented on the Equipment Master List or in individual calibration procedures. |  |  |  |  |
| --- | --- | --- | --- | --- |
| - Requirement for the use of suitably recognised traceable standards, and the requirement for maintaining certificates of traceability for all reference standards used in calibrations. - Requirement that there be a calibration procedure for each piece or type of equipment/instrument. Each procedure shall include: - Standard(s) to be used - Step by step calibration instructions. - Calibration tolerance. - Provisions for adjustments, if necessary. - Requirement for recording actual measurements found both before and after adjustment. - Format for the Calibration Record, including appropriate provisions for approval(s) of data. |  |  |  |  |
| - Provision for timely investigation of any Out of Tolerance (OOT) conditions. This must include: - An assessment of all production/work performed using the OOT equipment/ instrumentation since the previous acceptable calibration or verification/ accuracy check. - A review of that work to evaluate the possible impact upon product quality. - A documented rationale, approved by the Quality function, as to why product/work was or was not impacted. - Requirement that any OOT equipment/instrumentation be labelled Out of Service and not be used until it is adjusted and found to be within the required tolerances. - Following appropriate supervisory review and approval, calibration records shall be retained on a schedule commensurate with site record retention policies. Records should be retained at least as long as other records associated with product manufactured using the equipment/instrumentation in accordance with relevant SOPs |  |  |  |  |
| - An appropriate change control system, to cover: - Change in a calibration tolerance. - Change in calibration frequency. - Change in a calibration procedure. - Addition/deletion of equipment/instrumentation to/from the calibration programme. - Changes in location of equipment/instrumentation. - An appropriate procedure to tae account of errors of uncertainty in both the measured value of the instrument and the reference standards used for calibration - An appropriate procedure for ensuring all calibration certificates are filled in correctly and have the appropriate signatures and dates in place. - A method is in place to ensure that all involved in the calibration and verification processes are competent so to do, by documented assessment. - Internal staff - External staff |  |  |  |  |