Attachment 3a

Example Non-Conformance Sheet

| Deviation Report |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| [Company Name] [Company Address] Telephone Name Date Printed Page x of x |  |  |  |  |  |
| Department | XYZ | Deviation Date | dd/mm/yyyy | Deviation No. | 123 |
| See Calibration Cert No. | 123 | Previous Cal. Date | dd/mm/yyyy |  |  |
|  |  | Tag No. | AB-123-4567 |  |  |
| Deviation The frequency of this Instrument is greater than 123 days. It was calibration 123 days after its normal calibration rate. Errors based on (Before & After) Errors BEFORE Adjustment Rising Calibration Instrument Bench #3, Line 2, has exceeded the allowed tolerance. The indicated value was 0.00 (Lower Tol: 2.30, Upper Tol: 2.70) BEFORE Adjustment Rising Calibration Instrument Bench #3, Line 3, has exceeded the allowed tolerance. The indicated value was 0.00 (Lower Tol: 4.80, Upper Tol: 5.20) BEFORE Adjustment Rising Calibration Instrument Bench #3, Line 4, has exceeded the allowed tolerance. The indicated value was 0.00 (Lower Tol: 7.30, Upper Tol: 7.70) BEFORE Adjustment Rising Calibration Instrument Bench #3, Line 5, has exceeded the allowed tolerance. The indicated value was 0.00 (Lower Tol: 9.80, Upper Tol: 10.20) |  |  |  |  |  |
| Action Required: Replace transmitter |  |  | Raised By: First name Last name |  |  |
| Corrective Action Taken: Transmitter replaced (like for like). Accident & Incident 1234.567 raised. |  |  | See Calibration Cert. No: 123 |  |  |
| Results of Action: Pressure transmitter now functions within tolerance. |  |  |  |  |  |
| Signature: |  |  | Date: |  |  |