Attachment 12

Analytical Equipment Calibration Standard Operating Procedure

PRE PM

Standard Check Pre-PM reference:                          _ –Stdchk-__    __

Result: ________________

Performed By: __________________                      Date: ________________

Reviewed By: __________________                       Date: ______              ___

Comments:_____________________________________________________

_____________________________________________________

PM Performed By: __________________                      Date: _____________

PM Reviewed By: __________________                       Date: ______         ___

POST PM

Calibration Post-PM reference:                           _ –Cal-__    __

Result: ________________

Standard Check Post-PM reference:                         _ –Stdchk-__    __

Result: _____   _________

Precision Check reference:                         _ –Precision Check-__    __

Result: _____   _________

Performed By: __________________                      Date: ________________

Reviewed By: __________________                       Date: ______              ___

Comments:____________________________________________________

____________________________________________________

Checklist for Annual PM for xyz System

Serial Number of Instrument:_______________________

Chemistry Lab Id No:_________________________

Task to Be Performed						 Yes         No__

1. Replace Plunger and Face seals                                                                  _

2. Replace the seal wash assembly seals			                      _

3. Replace the Plungers						                      _

4. Replace the inlet check valves cartridges                                                   _

5. Replace the Reservoir Filters                                                                      _

6. Replace the in – line filter inserts                                                               _

7. Replace the sparge diffuser filters (If applicable)                                      _

8. Perform a static leak test to ensure that the solvent                                    _

delivery system is operating properly

Pass_____ Fail_____0

9. Press the Diag. screen key in the main screen, then record                       _

the total eluent delivered value displayed in the diagnostics

screen.: _______Litres.

10. Replace the Syringe.                                                                                  _

11. Replace the needle and the seal pack seals                                                _

12. Perform a purge and compression check to ensure that                            _

all fittings are tight.     Pass_____ Fail_____

13. Adjust the seal pack to optimise the needle port position.                        _

14. Enter the seal pack data values below: Seal Pack Data                             _

displayed value Top_____ Bottom_____ Width_____

15. Press the diag. screen key in the main screen, the record                          _

the injection count displayed in the diagnostic screen.

Number of injections: ____________

16. PM Label attached to HPLC.                                                                     _

(Note: Not all procedures apply to instruments, check N/A if the procedure does not apply)

Comments:____________________________________________________

____________________________________________________

Checklist for annual PM for xyz Absorbance detector

Note: Annual Preventative maintenance is not required for the abc refractive index

detector. All sections below can be marked N/A for this detector

Serial Number of Instrument:________________________

Chemistry Lab Id No:_________________________

Task to Be Performed						 Yes         No__

1. Replace the Lamp                                                                                        _

2. Replace the Battery                                                                                      _

3. Inspect the Fan                                                                                             _

4. Inspect the flow Cell                                                                                    _

5. Perform the diagnostic test: Pass: _____Fail:_____                                    _

6. PM Label attached to detector.                                                                     _

(Note: Not all procedures apply to instruments. Check N/A if the procedure

does not apply)

Comments:_____________________________________________________

_____________________________________________________

_____________________________________________________

PM Performed By: __________________                      Date: _____________

PM Reviewed By: __________________                       Date: ______         ___

PRECISION CHECK

Date:__________________

Method Reference: _________________ Method.

Calibration Reference: _____________________      __________

Standard Check reference:_________________       ___________

Precision Check reference: ________________- Precision Check-_____

Peak Area__

Injection 1							__________

Injection 2							__________

Injection 3							__________

Injection 4							__________

Injection 5							__________

Injection 6							__________

Average								__________

Standard Deviation (SD)					 _______ __

% Relative Standard Deviation (%RSD)			 __________

Specification							________ __

Result								_Pass / Fail_

Performed By: __________________                      Date: ________________

Reviewed By: __________________                      Date: ________________